2 Channel Bipolar Power Converter Chassis Production Test Procedure

(Inspection, Setup, Electrical, Functional, and Performance Testing)

(Rev 1.7, 6/12/26)

Chassis Model LT-EL-SR-PS-CRR-620x, x = 1 (S) or 2 (F)     Chassis S/N _________________

Date __________________

Prerequisites:

- Test conductor is familiar with power converter and chassis design.

- Test conductor has received training on this test procedure and is familiar with what the test results are expected to look like.

- Chassis undergoes EEI inspection upon receipt from vendor before AC power is applied.

- 208 VAC input section is draped when testing with AC power on and chassis top cover off.

- If any inspection or test step gives wrong result, abort inspection/test procedure, segregate chassis, and inform EEG group.

Note that the test procedure describes the steps needed to test one half of the power converter chassis. The test steps below should be carried out for both left (CH1) and right (CH2) halves of the power converter chassis. Place a check in the box after completing test step for both channels.

A. Visual Inspection (chassis top cover off, no AC power):

- Power entry module 208 VAC fast-on connectors are shrouded type.

- 208 VAC terminal block on Delta power supply has plastic slip-on finger guard in place.

- 208 VAC terminal block on Delta power supply middle terminal is connected to green ground.

- Delta power supply module connectors CN102 are well seated, and current share signal wire is connected between modules as needed.

- Six cables from Delta power supply to controller board are plugged into the right connectors.

- +24 V wiring polarity is correct for 12 AWG connections to power amplifier board (LT-EL-SR-PS-CRR-620x-11) terminals J3 and J7.

- +24 V wiring polarity is correct for 22 AWG connections to power amplifier power daughterboard (LT-EL-SR-PS-CRR-6301-21).

- +24 V wiring polarity is correct for 22 AWG connections to PWM-Regulator (LT-EL-SR-PS-CRR-620x-01) power daughterboard (LT-EL-SR-PS-CRR-6301-11).

- Power amplifier board electrolytic capacitors are installed correctly on PCB.

- Power amplifier gate drive daughterboards (LT-EL-SR-PS-CRR-6301-01) are installed.

- PWM-Regulator power daughterboard white testpoints and Controller board TP12 and TP14 do not read short to ground.

- Power amplifier +24 V DC bus does not measure shorted with an ohmmeter between terminals J3 and J7.

- Fused +24 V wiring polarity is correct for fan connector J1 on PWM-Regulator board.

- Fans are facing the correct way.

- Wiring polarity is correct for voltage feedback twisted pair connection between power amplifier board J5 and PWM-Regulator board J10.

- PWM-Regulator board jumpers shunts are installed correctly according to the following table.

| Ref. Designator | Installed | Not Installed | Function |
| --- | --- | --- | --- |
| J15 |  |  | HALL1A Current feedback |
| J9 |  |  | HALL2A Current feedback |
| J17 |  |  | HALL1B Current feedback |
| J12 |  |  | HALL2B Current feedback |
| J14 |  |  | HALL1A Current feedback gain bypass |
| J2 |  |  | HALL2A Current feedback gain bypass |
| J16 |  |  | HALL1B Current feedback gain bypass |
| J7 |  |  | HALL2B Current feedback gain bypass |
| J4 |  |  | 1 MHz clock |
| J22 |  |  | 1 MHz clock termination |
| J18 |  |  | -2.5 V gain bypass |
| J8 |  |  | Voltage feedback |
| J3 |  |  | PID voltage loop gain bypass |

- PID output 3-position screw terminal block J13 has jumper short installed.

- Ethernet cable is installed.

- Ribbon cable installed between PWM-Regulator board J11 and power amplifier board J6.

- Ribbon cable installed between PWM-Regulator board J6 and power amplifier board J2.

- PCB mounting screws are tight.

- 12 AWG wire connections are tight.

- 10 A fuses are installed in 208 VAC power entry module.

- 1-wire sensors are attached to heatsinks and wired correctly to screw terminal block J2 on controller board (LT-EL-SR-PS-CRR-620x-21).

- On power daughterboard (LT-EL-SR-PS-CRR-6301-11) for each PWM-Regulator board, U1 has RS6-2415D part number.

- Verify that both heatsink pieces for each channel have a 10 kΩ connection to chassis.

B. Microcontroller Programming (chassis top cover off, no AC power)

- Burn bootloader through ICSP connector with Arduino programmer (Tools->Burn Bootloader).

- Program Microcontroller through ICSP connector with Arduino programmer (Sketch->Upload Using Programmer).

- On separate PC, add file “config.txt” to SD card with information in the format below (a template config.txt file is available in /home/dib/2ch/):

  - IP Address: 192.168.0.y where y is the serial number of the unit.

  - MAC Address: 02,00,00,62,0x,y where x is either 1 or 2 for slow and fast, and y is the serial number of the unit in hexadecimal (base 16) representation.

  - 1-Wire Sensor Left: 28,00,00,00,00,00,00,00

  - 1-Wire Sensor Right: 28,00,00,00,00,00,00,00

  - HALL sensor gain: 1.02, 1.02

  - Model.Serial Number: 620x00y where x is either 1 or 2 for slow and fast, and y is the serial number of the unit.

- Insert SD card into SD module on controller board.

C. Initial Checkout (chassis top cover off, AC power ON)

Controller Board Checkout

- Remove fuses F1 and F3 from both Power Amplifier Power daughterboards (LT-EL-SR-PS-CRR-6301-21).

- Remove SiCFET gate drive daughterboards from power amplifier boards.

- Remove fuse F1 from controller board (LT-EL-SR-PS-CRR-620x-21).

- With 208 VAC section properly draped, apply AC power to chassis, and verify with a DMM that there is +5 V power at controller board fuse F1 socket. Use TP9 or TP10 as circuit common reference. If DMM reading shows -5 V, stop test, and correct wiring error.

- Turn off AC power, install fuse F1, and turn AC power back on.

- Verify 3.3 V at TP11.

- Verify that amber LED D7 (heartbeat) on controller board is blinking.

- In separate terminal window on the workstation computer, run the housekeeping data monitor program by typing

sudo python3 hyperUDPdib.py 192.168.0.y where y is the serial number of the unit under test.

- Connect USB-TTL-Serial adapter to controller board J18 header. In a separate terminal window, run serial terminal program by typing

sudo python3 hypertermdib.py /dev/ttyUSB x  where x is usually 0 but could be 1 or 2.

Reset controller board or ac cycle power converter, and observe startup text in terminal.

Type ctrl-C to stop the terminal program.

Get 1-wire temperature sensor ID codes, and update config.txt file on SD card with ID codes by typing

sudo python3 pushSDdib1WireIDs.py 192.168.0.y  where y is the serial number of the unit under test.

Then select lines 3 and 4 to edit left and right sensor IDs Type ctrl-C to stop the terminal program.

PWM Board Checkout (gate drive boards NOT installed on power amplifier)

- On PWM-Regulator boards, temporarily move jumper shunt from header J8 to header J3.

- With bench channel controller board DB25 cables connected to chassis rear panel DB25 connectors, turn on power converter channels using the channel controller program. Do this in a separate terminal window by typing

sudo python3 alsu4chan.py /dev/ttyUSB1.

Note that to turn a channel on, both ON2 command bits from the PSC or channel controller must be set high first.

- Verify -5 V, -15 V, +5 V, and +15 V power supply levels at testpoints on power daughterboard (LT-EL-SR-PS-CRR-6301-11) for each PWM-Regulator board. Readings should be within ±0.2 V of nominal.

- Verify that heatsink fan turns on for several seconds then turns off.

- With an oscilloscope verify the following signals on the PWM-Regulator boards:

  - TP7, 250 kHz clock.

  - Triangle waves at TP22 and TP20. Peak amplitude should be ~4.5 V.

  - Observe PWM signals at C6, C9, C10, C25, C83, C88, C91, and C92. Verify that each of the eight waveforms has a fast rising edge and a slow falling edge.

- Turn channel off, then on again. Verify that PWM1- (J19-14, purple) and PWM2- (J19-18, white) go from high to low after about 2 s. Do not put hand near fan when turning on.

- Turn channel off, then on again. Verify that PARK1 (J19-12, green) goes from high to low after about 4 s. Do not put hand near fan when turning on.

- Connect signal source to the bench channel controller board setpoint input for each channel labeled as CHx_Regout (rear panel Dsub connectors pins 1 and 14). Insert gate drive board extender board into power amplifier board. In terminal window on workstation, type

sudo  python3 sweep1.py.

Using two differential oscilloscope probes, observe complementary PWM signals at testpoint header pins (PWM_IN, GND) on gate drive board extender. Measure four times, twice for each gate drive board position on the power amplifier board.

- Turn off power converter channels. Turn off AC power to chassis.

Power Amplifier Board Checkout

- Measure Rgs resistance at each SiCFET between pins 3 and 4. Verify 10 kΩ reading.

- Install both gate drive boards into power amplifier board.

- Install fuse F3 into power amplifier board power daughterboard.

- Connect a pair of clip leads across unpopulated F1 fuse holder terminals on the power amplifier power daughterboard.

- Apply AC power to chassis. Turn on the channel and measure current with an ammeter.

- Verify that 15 V current draw is approximately 0.3-0.4 A.

- Install fuse F1 into power daughterboards (LT-EL-SR-PS-CRR-6301-21).

- On PWM-Regulator boards, move jumper shunt from header J3 back to header J8.

- Turn on chassis AC power, and turn on each power converter channel. Verify +15 V power supply level at testpoint on power daughterboard (LT-EL-SR-PS-CRR-6301-21) for each power amplifier board.

- Connect signal source to the bench channel controller board setpoint input for each channel labeled as CHx_Regout (rear panel Dsub connectors pins 1 and 14).

In terminal window on workstation, type

sudo python3 sweep2.py.

With one differential oscilloscope probe connected across H-bridge outputs TP5/TP6 and another differential oscilloscope probe connected across H-bridge outputs TP7/TP8, verify proper output pulses as control input is varied. TP5 and TP7 are positive. TP6 and TP8 are negative.

D. External Tests (chassis top cover on, no load):

- Ignd readback ohmmeter verification. With both channels turned off, and using connector labeled CHx_Ignd on bench channel controller board, verify:

  - Ch1 minus output terminal to rear panel jumper block terminal 2 measures 100 Ω

  - Ch1 minus output terminal to rear panel Dsub connector pin 4 (CHx_Ignd connector) measures short

  - Rear panel terminal block terminal 1 measures short to chassis

  - Rear panel Dsub connector pin 17 (CHx_Ignd connector) measures short to chassis

- On bench channel controller board, verify that ON readback signal lights up high when channel is turned on (rear panel DB25 connector pins 20,24).

- Turn on channel with ON2 signal disabled to produce an ON fault. On bench channel controller board, verify that Fault1 signal lights up when channel is turned on (rear panel DB25 connector pins 21,24).

- On bench channel controller board, verify that Fault2/Heartbeat signal is pulsing once per second (rear panel DB25 connector pins 22,24).

- Verify that power converter output reads ±9.5 V with a ±5 V input to the power converter. If power converter output is larger than ±9.7 V, stop test.

- Verify that power converter output reads ±19 V with a ±10 V input to the power converter.

- At bench channel controller board connector CHx_Vout (rear panel Dsub connector pins 3,16), verify that power converter Vout readback signal reads -/+10 V ± 0.15 V with a ±10 V input to the power converter. Note that Vout readback is inverted from actual power converter output voltage.

- At bench channel controller board connector CHx_Spare (rear panel Dsub connector pins 5,18), verify that 24 V bus readback signal shows 4.8 V ± 0.1 V.

- Observe heatsink temperature readbacks in UDP housekeeping data. Put your finger on the 1-wire sensor on the CH1 heatsink. Verify that channel 1 temperature increases. If channel 2 temperature increases, swap the 1-wire temperature sensor ID codes in the config.txt file on the SD card by typing the following command:

sudo python3 swap1wireIDs.py 192.168.0.y  where y is the serial number of the unit under test

E. Performance Tests (chassis top cover on, with loads)

In terminal window on workstation, type sudo python3 bpc_test.py to run python program that will guide you through performance tests.

For steps 68-71, put OVC switch on load box in DOWN position.

- Noise

  - Normal mode switching noise and spectrum

  - Common mode switching noise and spectrum

  - Low frequency spectrum (need 10 kHz low pass filter)

- Step response and step delay

- Energy extraction (Load with 20 mH inductor)

- Demonstrate simultaneous max V and I

  - Measured  ____________ V and ______________ A CH1

  - Measured  ____________ V and ______________ A CH2

- OVC Interlock (OVC switch on load box in UP position).

  - +OVC Verified CH1

  - -OVC Verified CH1

  - +OVC Verified CH2

  - -OVC Verified CH2

- Email data to offline host by typing sudo ./tar1.sh (optional).

End of Test

Notes: _____________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

Print out all waveform plots generated from performance tests for channels 1 and 2, and include with this test procedure.

Test Conductor _____________________________________   Date _______________________

Performance Test measurements reviewed and approved by NSLS-II EEG group.

Engineer _________________________________   Date______________________